## Maufacturer Med. Device
Evaluation Criteria for Manufacturers of Medical Devices, Surgical Disposibles and Sutures for Government MCC 2025-26
| Name of the firm | [5 empty] | Pak Disposables Syringes Gadoon Sawabi
|  | Product General Information |  |  |  | Technical Evaluation Matrix
| S.No |  |  |  |  | Factory Technical Evaluation Parameter | [7 empty] | Factory Evaluated Score | Product technical Evaluation Parameters | [5 empty] | Product Evaluated Score | Total Technical Score
[6 empty] | Documents Based Factory Score |  |  |  | Evaluation Visit Score
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
[6 empty] | Valid ISO 14001 certificate of the facility where the quoted product is manufactured, issued by PNAC accredited body (duly attested by senior executive of the firm). Online verification link shall be provided. | Valid ISO 13485 certificate of the facility where the quoted product is manufactured, (duly attested by senior executive of the firm). Online verification link shall be provided. | Valid calibration certificates for equipment / instruments used in the factory for Measuring, weighing, Assay/ Analysis of raw material, in-process material and finished products for the manufacturing of the quoted products. (Valid Calibration Certificates attested by Quality head of the firm). | Valid documents of the Federal Board of Revenue (FBR) showing the total financial turnover of the firm for the last year i.e., FY 2023-24 or latest. Maximum 6 marks shall be awarded in the following manner: Financial turnover of PKR 100 to 500 million - 2 marks. Financial turnover of more than PKR 500 million and upto 1000 million - 4 marks. Financial turnover of more than PKR 1000 million - 6 marks (The document shall be attested by a Senior executive of the firm) | Adherence to Good Storage practices (GSP) for Raw material, In-process and Finished Goods. (as evaluated at the time of inspection by the MCC expert/s). Non adherence to GSP shall lead to disqualification of the firm. | Adherence to Current Good Manufacturing Practices (cGMP) in line with the DRAP regulations. (to be evaluated by the MCC expert/s at the time of inspection, Noncompliance to cGMP shall lead to disqualification of the relevant section or firm) | Availability of, Functional and validated HVAC, with all relevant equipment, testing, and logs. (As evaluated by the MCC expert/s at the time of inspection). Non-availability or non-functionality of the HVAC system and/or testing, and logs shall lead to Disqualification of the relevant section / firm. | Adequate availability of qualified & relevant Human Resource as per the requirements laid down in DRAP regulations. (Certified by the senior executive of the firm & evaluated by MCC expert/s at the time of inspection, Non-availability shall lead to disqualification of the section/s or firm). |  | Goods Declaration certificate of imported raw material of the quoted item/s from Pakistan Customs, coupled with valid airway bill or Bill of Lading for the quoted item/s, not older than 24 months on the cutoff date for submission of bids. In cases where Raw materials are acquired from Local sources valid invoice (s) not older than 24 months shall be considered. In case of purchases through third party importers a valid trail/link/DRAP clearance NOC between the principal manufacturer and the importer firm shall be established with the firm offering the product to Govt. MCC (Certificate Duly attested by Senior Executive of the firm) | Certificate of Analysis of raw material from the Principal Manufacturer as mentioned in the goods declaration (GD) provided in column 14, duly attested by the senior executive of the firm. In case of Non-provision of matching GD the marks for CoA will not be awarded. | Tender Approvals (not older than twelve 12 months) from Tertiary care Govt. Hospitals, Health related Govt. projects and/ or JCI accredited private hospitals of Pakistan. Marks shall be awarded in the following manner: 02 Tender approvals- 01 mark 04 Tender approvals- 02 marks 06 Tender approvals- 03 marks 08 Tender approvals- 04 marks 10 or more Tender approvals- 05 marks Note. Tender approval means award of contract(s) for the quoted product(s) with the same brand name and specifications / strength / dosage form. Moreover, the approval(s) shall be duly attested by the concerned procuring entities and accompanied with purchase orders, delivery challan, invoice/warranty and cheque from the concerned procuring entity/purchasing agency/ies, etc. The bank statement showing the respective transaction against the cheque received shall also be submitted with each purchase order. | Valid WHO prequalification and / or valid product registration in SRA country(ies) / and / or valid free sale certificate issued by regulatory body of any SRA country(ies) In cases where the validity period is not explicitly mentioned on the above certificate, the certificate shall be considered valid only if it was issued within the last five (05) years from the date of bid submission. 2 marks for each certification, up to a maximum of 06 marks. Certificates on company's own letter heads shall not be acceptable. (copies of relevant certificates duly attested by the senior executive of the firm) | Samples evaluation by DTL (Failure to comply with relevant standards shall lead to Disqualification of the quoted products) | Physical examination of the quoted item/s by the MCC expert/s. Rejection of the quoted item/s by the MCC expert/s shall lead to disqualification of the said item/s.
|  | Ref. No. of item in MCC Formulary | Generic Name of Item | Size & Guage of MedicaL Device | Trade Name | 3 | 5 | 5 | 6 | 3 | 3 | 2 | 2 | 29 | 5 | 5 | 5 | 6 | 10 | 10 | 41 | 70
1 | 1016 | Disposable Auto Disable Syringe (Blister packing) sterile | 3 ml | PDS AD Syringe |  | 3 | 0 | 5 | 0 | The firm was inspected by the inspection team as per technical evaluation criteria, and the following observations were made; Compromised Good Manufacturing Practices (cGMP): 1. Production process incomplete with gaps in implementation, creating potential risks to product consistency. 2. SOPs for production flow and validation are not fully implemented. 3. Overall Hygiene was compromised in quoted sections. 4. The Calibration of multiple equipment in the QC and microbiology section were expired. 5. No record of recent validation/qualification or preventive maintenance record of the equipment were provided during the inspection. Impaired environmental control and HVAC System: 6. The HVAC (16+2 tons) system was found non-functional during inspection. 7. The SOPs for entering the temperature and humidity maintenance logs were missing. 8. Open ventilation and room cooler were observed in the production section during active production of disposable syringes 9. Sterility and pharmaceutical hygiene were compromised due to inadequate air handling and environmental control. 10. Fake entries in the log sheets or no temperature/humidity mapping, was maintained. Compromised Good Storage Practices (GSP) 11. Raw materials WERE stored in one location without segregation for approved, rejected, or quarantined lots. 12. The RMS was an open area with no environmental controls. 13. The Packing material, active material and rejected material were lying in a single section. 14. The overall hygiene in the so-called RMS was very poor and unhygienic. 15. No dedusting on receiving as performed for the received raw materials. 16. No GRN (Goods Receipt Note), QC, or QA rejection protocols observed. 17. No Segregated areas with environmental monitoring systems are were observed during inspection essential to prevent cross-contamination and ensure traceability. Quality Control (QC) / Quality Assurance (QA) / GLP 1. No qualified staff was available in the QC or Quality assurance to perform the requisite tests at the time of inspection. 2. Microbiology lab equipped but non-operational due to lack of qualified microbiologist. 3. Validation and in-process testing procedures absent. 4. Documentation incomplete or missing for analytical testing. In view of the above findings, the firm is NOT-RECOMMENDED. |  |  |  | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8
2 | 1017 | Disposable Auto Disable Syringe (Blister packing) sterile | 5 ml | PDS AD Syringe |  | 3 | 0 | 5 | 0 |  |  |  |  | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8
3 | 1921 | Disposable Syringe Ordinary (Blister packing) sterile | 10ml | PDS Syringe |  | 3 | 0 | 5 | 0 |  |  |  |  | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8
4 | 1112 | I/V Fluid administration set | Sterile and pyrogen free, minimum 150cm tube length, blister pack | PDS I.V Set |  | 3 | 0 | 5 | 0 |  |  |  |  | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8
GRC Decision
The GRC agreed with the findings and recommendations of the inspection report and disposed of the appeals in accordance with the said terms.