## Cotton Related Goods Manufactur
Evaluation Criteria for Manufacturers of Cotton & Related Goods for Government MCC 2025-26
| Name of Firm |  | 63 HAFIZ PHARMA INDUSTRY KAMOKE
S. No. | General Product Information |  |  |  | Technical Evaluation Matrix
[5 empty] | Factory Technical Evaluation Parameters | [9 empty] | Factory Evaluated Score | Product Evaluation Parameters |  |  | Product Evaluated Score | Total Technical Score
[5 empty] | Documents Based Factory Score | [5 empty] | Evaluation visit Score
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
[5 empty] | Valid ISO 14001 certificate of the facility where the quoted product is manufactured, issued by PNAC accredited body (duly attested by senior executive of the firm). Online verification link shall be provided. | Valid ISO 45001 certificate of the facility where the quoted product is manufactured, issued by PNAC accredited body (duly attested by senior executive of the firm). Online verification link shall be provided. | Valid ISO 9001 certificate of the facility where the quoted product is manufactured, issued by PNAC accredited body. (duly attested by senior executive of the firm). Online verification link shall be provided. | Valid ISO 13485 certificate of the facility where the quoted product is manufactured, (duly attested by senior executive of the firm). Online verification link shall be provided. | Valid calibration certificates for equipment / instruments used in the factory for Measuring, weighing, Assay/ Analysis of raw material, in-process material and finished products for the manufacturing of the quoted products. (Valid Calibration Certificates attested by Quality head of the firm). | Valid documents of the Federal Board of Revenue (FBR) showing the total financial turnover of the firm for the last year i.e FY 2023-24 or latest. A minimum turnover of PKR 100 million is required for award of marks in this parameter. (The document shall be attested by a Senior executive of the firm) | Functional and effective Airconditioning & Ventilation System as per the requirements laid down by DRAP (Evaluated by the MCC expert/s at the time of inspection, Non functionality of the Air Conditioning & Ventilation system in specified section shall lead to disqualification of the section or firm) | Adequate availability of equipment / instruments in QC labs performing relevant official tests as well as compliance to Good laboratory practices (GLP) in all Labs and Current Good Manufacturing Practices (cGMP) throughout the production facility. (Evaluated by the MCC expert/s at the time of inspection, Non availability of adequate and appropriate equipment / instruments and non-compliance to GLP , cGMP shall lead to disqualification of the relevant section or firm) | Appropriate storage of raw material, in process and finished goods with compliance to Good storage practices (GSP) (To be evaluated by the MCC expert/s at the time of inspection, Non compliance to GSP shall lead to disqualification of the relevant section or firm) | Adequate availability of qualified & relevant Human Resource as per the requirements laid down in DRAP regulations. (Certified by the senior executive of the firm & evaluated by MCC expert/s at the time of inspection, Non-availability shall lead to disqualification of the section/s or firm). |  | Tender Approvals (not older than twelve 12 months) from Tertiary care Govt. Hospitals, Health related Govt. projects and/ or JCI accredited private hospitals of Pakistan. Marks shall be awarded in the following manner: 02 Tender approvals- 01 mark 04 Tender approvals- 02 marks 06 Tender approvals- 03 marks 08 Tender approvals- 04 marks 10 or more Tender approvals- 05 marks Note. Tender approval means award of contract(s) for the quoted product(s) with the same brand name and specifications / strength / dosage form. Moreover, the approval(s) shall be duly attested by the concerned procuring entities and accompanied with purchase orders, delivery challan, invoice/warranty and cheque from the concerned procuring entity/purchasing agency/ies, etc. The bank statement showing the respective transaction against the cheque received shall also be submitted with each purchase order. | Physical examination of the quoted item/s by the MCC expert/s. Rejection of the quoted item/s by the MCC expert/s shall lead to disqualification of the said item/s. | Samples evaluation by DTL (Failure to comply with relevant standards shall lead to Disqualification of the quoted products)
| Ref. No. of item in MCC Formulary | Generic Name of Item | Sizes and specifications | Trade Name | 3 | 3 | 3 | 4 | 5 | 3 | 6 | 6 | 6 | 6 | 45 | 5 | 10 | 10 | 25 | 70
1 | 967 | Cotton (Surgical) Corded BPC | 200 GM | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 | The inspection team at the time of inspection observed that, POOR ADHERANCE TO cGMP 1. Only one registered pharmacist was present on-site at the time of inspection, which is insufficient for ensuring adequate supervision across all manufacturing, quality control, and storage areas as per GMP and regulatory staffing requirements. 2. Materials stored in the RMS and packaging storage area were not labeled, violating essential GSP principles related to material identification, status indication, and traceability. 3. No qualified microbiologist was available at the time of inspection, indicating a deficiency in the facility’s quality control capability, particularly for microbiological testing and environmental monitoring POOR GOOD STORAGE PRACTICES (GSP) 4. No hygrometers or humidity monitoring devices were installed in the Raw Material Store (RMS), in-process areas, or Finished Goods Store (FGS), preventing real-time monitoring of environmental conditions critical for material stability. 5. No temperature or relative humidity records were maintained in any of the storage or production areas, indicating a serious lapse in environmental monitoring and documentation control. 6. Finished goods were stored directly on the floor without the use of pallets or storage racks, increasing the risk of contamination and product damage due to poor storage practices. In view of the above the firm is NOT RECOMMENDED and marks in importer’s evaluation parameters for the quoted items may be awarded |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
2 | 968 | Cotton (Surgical) Corded BPC | 100 GM | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 |  |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
3 | 969 | Cotton Bandages (Surgical) B.P Type II | 6.5cm x 4 m | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 |  |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
4 | 970 | Cotton Bandages (Surgical) B.P Type II | 7.5cm x 4m | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 |  |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
5 | 971 | Cotton Bandages (Surgical) B.P Type II | 10cm x 4m | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 |  |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
6 | 972 | Cotton Bandages (Surgical) B.P Type II | 15cm x 4m | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 |  |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
7 | 979 | Cotton Crepe Bandage BPC | 7.5cm x 4.5m | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 |  |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
8 | 980 | Cotton Crepe Bandage BPC | 10cm x 4.5m | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 |  |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
9 | 981 | Cotton Crepe Bandage BPC | 15cm x 4.5m | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 |  |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
10 | 1092 | Gauze Cloth Roll Packing | 100cm x 20m | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 |  |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
11 | 1093 | Gauze Cloth Roll Packing | 100cm x 40m | Surgeon | 3 | 3 | 3 | 4 | 5 | 3 |  |  |  |  | 21 | 0 | 10 | 10 | 20 | 41
GRC Decision
The GRC agreed with the findings and recommendations of the inspection report and disposed of the appeals in accordance with the said terms.